## Input
Audiogram
| Frequency [Hz] | Hearing level [dB]
| 125 | 40
| 250 | 60
| 500 | 75
750
| 1000 | 110
1500
| 2000 | 120
| 4000 | 105
6000
8000
Geometric data
Threshold level [dB] |  |  | 65
Insertion angle [°] |  |  | 602
Array insertion depth [mm] |  |  | 27.8
Type of electrode array |  |  | FLEXsoft
Surgical approach |  |  | RW
1
Estimated number of electrodes in acoustic region |  |  | 1

## Calculations (hidden)
1. Cut-off frequency for acoustic region determined from audiogram
Index: If both HL at this and previous frequency are <= threshold level and the same did hold for the previous cell --> give index number;
Audiogram (copy of input) |  |  | else if HL > threshold and previous hearing level <= threshold and previous cell has an index --> give index number; else 0
Frequency [Hz] |  | Hearing level [dB] | Index | HL <= TL | Last entry
125 |  | 0 | 1
125 |  | 40 | 2 | 1 | 2
250 |  | 60 | 3 | 1 | 3
500 |  | 75 | 4 | 0 | 4
750 |  | 92.5 | 0 | 0 | 0
1000 |  | 110 | 0 | 0 | 6
1500 |  | 115 | 0 | 0 | 0
2000 |  | 120 | 0 | 0 | 8
4000 |  | 105 | 0 | 0 | 9
6000 |  | 52.5 | 0 | 1 | 0
8000 |  | 0 | 0 | 1 | 0
|  |  | Index to use: | 4 | 9
Interpolate between these points
|  |  | x1 | 250
Cut-off frequency (Fc) |  |  | y1 | 60
333.3333 |  |  | x2 | 500
|  |  | y2 | 75
2. Values to adapt insertion depth
If cochleostomy is used add value from input [mm] |  |  |  | 0
Distance tip to 1st electrode [mm] |  |  |  | 1.2
Degree shift for uncertainty range [°] |  |  |  | 24
3. Calculation for array
Description |  |  |  | Value | Value angle + x ° | Value angle - x °
Insertion depth E1 at RW 0 ° |  |  |  | 26.6
Relative length OC at cut-off freq |  |  |  | 0.78014
Angle at cut-off freq for OC |  |  |  | 544.9769 | 568.9769 | 520.9769
Relative length OC at cut-off freq |  |  |  |  | 0.79748 | 0.76196
Rel length OC at insertion angle |  |  |  | 0.82001
Relative length OC at middle RW (Stakhovskaya; middle of RW is 1.5 degree for the beginning of OC) |  |  |  | 0.053295
OC length + geometry |  |  |  | 34.3899
Acoustic region |  |  |  | 26.8291 | 27.4252 | 26.2037
Length of OC in acoustic region [mm] |  |  |  | 1.3709 | 0.77482 | 1.9963
Number of electrodes in acoustic region |  |  |  | 1.5712 | 1.3228 | 1.8318
Data to plot uncertainty area
1.3228 | -0.1
1.3228 | 0.1
1.3725 | -0.1
1.3725 | 0.1
1.4222 | -0.1
1.4222 | 0.1
1.4719 | -0.1
1.4719 | 0.1
1.5215 | -0.1
1.5215 | 0.1
1.5712 | -0.1
1.5712 | 0.1
1.6233 | -0.1
1.6233 | 0.1
1.6755 | -0.1
1.6755 | 0.1
1.7276 | -0.1
1.7276 | 0.1
1.7797 | -0.1
1.7797 | 0.1
1.8318 | -0.1
1.8318 | 0.1

## Electrode arrays (hidden)
MED-EL electrode array | Distance between electrodes [mm] | Length of tip [mm] | Surgical approach
FLEXsoft | 2.4 | 1.2 | RW
FLEX28 | 2.1 | 1.2 | Cochleostomy
FLEX26 | 1.9 | 1
FLEX24 | 1.9 | 1
FLEX20 | 1.4 | 1
Standard | 2.4 | 1.2
Medium | 1.9 | 1.2
Compressed | 1.1 | 1
Value of selected array | 2.4 | 1.2

## Reference
Natural Electro Acoustic Stimulation Fitting in Cochlear Implant Patients with Hearing preservation
An increased number of people with low frequency hearing are now being considered for cochlear implantation. Nowadays, it is common to preserve the residual hearing with deeply inserted electrodes (Skarzynski, Lorens, 2010; Skarzynski et al., 2011; Helbig et al., 2011; Usami et al., 2011; Bruce et al., 2011, 2014; Rogers et al., 2012; Mick et al., 2013; Nordfalk et al., 2016; O'Connell et al., 2017; Moteki et al., 2018). As such, there is an increased probability that the electrode array could be positioned within the acoustic region of the cochlea and a certain amount of overlap between electric and acoustic stimulation within the cochlea could be expected. In order to optimize the overlap between the acoustic hearing and the electrical stimulation the following EAS Fitting method can be applied: 1. Calculate how many electrodes are in the acoustic region (postop CT required) 2. Consider switching off electrodes in the acoustic region based on the results provided by the calculations above 3. Apply the EAS fitting method in the MAESTRO fitting software The first point is achieved via the Excel tool. For the fitting purposes the acoustic region is defined at 65dB HL from the postoperative audiogram. The precision of this tool is +24 °, which accounts for + 0.84mm of the electrode array length. Please mind that this uncertainty is not accounted for the estimated number of electrodes in the acoustic region, but it is visualized in the graph. The necessary input is hearing levels at characteristic frequencies, electrode insertion angle, electrode insertion depth, type of array, type of surgery (RW/cochleostomy – default 1mm from RW). All the input parameters are highlighted in turquoise color. The next part of the fitting is performed in the MAESTRO clinical software (please always use the latest version). In the MAESTRO software you have the option to switch the identified electrodes off and the non-acoustic region electrode are activated with the frequency range starting from the cut off frequency (typically at 65dBHL from the unaided audiogram). Make sure you select FS4 speech coding strategy. Please check if the insertion angle and depth in the input area are correct. The previous research showed that complementary fine structure as implemented in the FS4 strategy, improves hearing of EAS subjects especially in noise. Having this, the EAS fitting continues as a regular EAS fitting (MED-EL EAS fitting manual).
References
Greenwood D.D. (1990) A cochlear frequency-position function for several spieces - 29 years later. Journal of the Acocustic Society of America 87(6):2592-2605.
Kawano, A., Seldon, H.L., Clark, G.M. (1996) Computer-aided three-dimensional reconstruction in human cochlear maps: measurement of the lengths of organ of Corti, outer wall, inner wall, and Rosenthal´s canal. Ann. Otol. Rhinol. Laryngol. 105(9):701-709
Polak M, Lorens A, Furmanek M, Skarzyński H. (2020) Electrode estimation in the acoustic region of the human Cochlea. Acta Otolaryngol. 30:1-10. doi: 10.1080/00016489.2020.1736337.
Stakhovskaya, O., Sridhar, D., Bonham, B.H., Leake, P.A. (2007) Frequency map for the human cochlear spiral ganglion: implications for cochlear implants. Journal of the association for research in otolaryngology 8:220-233